## Table S1
Table 1S. Comparison of major and trace element concentrations in reference samples analyzed using X-Ray Fluorescence spectrometry and Inductively Coupled Plasma-Mass Spectrometry, as well as detection limits for both analytical methods. The number of replicate analyses for XRF and ICP-MS methods is the same.
|  | BE-N |  |  | BHVO-1
|  | Recomm. | Measured | Relative | Recomm. | Measured | Relative | Detection
|  |  | Mean Values | Error (%) |  | Mean Values | Error (%) | limits
|  |  | (n=32) |  |  | (n=15)
X-Ray Fluorescence spectrometry (XRF) | (wt%)
| SiO2 | 38.4769 | 38.57 | -0.24189 | 49.94 | 49.57 | 0.74089 | 0.05
| TiO2 | 2.6289 | 2.6207 | 0.31417 | 2.71 | 2.75 | -1.476 | 0.01
| Al2O3 | 10.143 | 9.8 | 3.3817 | 13.8 | 14.02 | -1.5942 | 0.05
| Fe2O3 | 12.9331 | 12.79 | 1.1063 | 12.23 | 12.5 | -2.2077 | 0.1
| MnO | 0.20145 | 0.18645 | 7.4462 | 0.17 | 0.17 | 0 | 0.05
| MgO | 13.2453 | 13.61 | -2.7532 | 7.23 | 6.9 | 4.5643 | 0.01
| CaO | 13.9705 | 13.73 | 1.7218 | 11.4 | 11.52 | -1.0526 | 0.04
| Na2O | 3.2031 | 3.28 | -2.4023 | 2.26 | 2.34 | -3.5398 | 0.01
| K2O | 1.4001 | 1.36 | 2.8625 | 0.52 | 0.5 | 3.8462 | 0.01
| P2O5 | 1.0576 | 1 | 5.4474 | 0.27 | 0.25 | 7.4074 | 0.01
(ppm)
| Zn | 120 | 116 | 3.3333 | 105 | 98 | 6.6667 | 2
| Cu | 72 | 72.5759 | -0.79981 | 136 | 140 | -2.9412 | 3
| Sc | 22 | 21.4931 | 2.3041 | 31.8 | 33 | -3.7736 | 3
| Ga | 17 | 16.2586 | 4.3611 | 21 | 22 | -4.7619 | 3
| Ni | 267 | 260 | 2.6217 | 121 | 124 | -2.4793 | 2
| Co | 60 | 63 | -5 | 45 | 47 | -4.4444 | 2
| Cr | 360 | 352 | 2.2222 | 289 | 298 | -3.1142 | 2
| V | 235 | 231 | 1.7021 | 317 | 312 | 1.5773 | 2
| Rb | 47 | 46.8889 | 0.23641 | 11 | 10 | 9.0909 | 1
| Ba | 1025 | 1001 | 2.3415 | 139 | 146 | -5.036 | 3
| Pb | 4 | n.d. |  | 2.6 | n.d. |  | 5
| Sr | 1370 | 1360 | 0.72993 | 403 | 408 | -1.2407 | 2
| Zr | 260 | 268 | -3.0769 | 179 | 173 | 3.352 | 2
| Y | 30 | 28.4 | 5.3333 | 27.6 | 28 | -1.4493 | 1
| La | 82 | 79 | 3.6585 | 15.8 | 18 | -13.9241 | 5
| Ce | 152 | 160 | -5.2632 | 39 | 41 | -5.1282 | 5
| Nd | 67 | 65 | 2.9851 | 25.2 | 24 | 4.7619 | 3
| Nb | 105 | 105 | 0 | 19 | 18 | 5.2632 | 1
| Th | 10.4 | 11 | -5.7692 | 1.08 | 1 | 7.4074 | 1
Inductively Coupled Plasma-Mass Spectrometry (ICP-MS) | (ppm)
| Rb |  |  |  | 11 | 10.8 | 1.8182 | 0.02
| Sr |  |  |  | 403 | 410 | -1.737 | 0.02
| Y |  |  |  | 27.6 | 28.2 | -2.1739 | 0.003
| Zr |  |  |  | 179 | 190 | -6.1453 | 0.02
| La |  |  |  | 15.8 | 16.01 | -1.3291 | 0.01
| Ce |  |  |  | 39 | 37.8 | 3.0769 | 0.002
| Pr |  |  |  | 5.7 | 5.57 | 2.2807 | 0.004
| Nd |  |  |  | 25.2 | 26.12 | -3.6508 | 0.002
| Sm |  |  |  | 6.2 | 6.53 | -5.3226 | 0.006
| Eu |  |  |  | 2.06 | 2.1 | -1.9417 | 0.002
| Gd |  |  |  | 6.4 | 6.25 | 2.3438 | 0.001
| Tb |  |  |  | 0.96 | 0.925 | 3.6458 | 0.002
| Dy |  |  |  | 5.2 | 5.42 | -4.2308 | 0.001
| Ho |  |  |  | 0.99 | 1.05 | -6.0606 | 0.002
| Er |  |  |  | 2.4 | 2.57 | -7.0833 | 0.001
| Tm |  |  |  | 0.33 | 0.318 | 3.6364 | 0.003
| Yb |  |  |  | 2.02 | 2.16 | -6.9307 | 0.001
| Lu |  |  |  | 0.291 | 0.309 | -6.1856 | 0.005
| Nb |  |  |  | 19 | 18.48 | 2.7368 | 0.001
| Hf |  |  |  | 4.38 | 4.29 | 2.0548 | 0.001
| Ta |  |  |  | 1.23 | 1.33 | -8.1301 | 0.005
| Th |  |  |  | 1.08 | 1.13 | -4.6296 | 0.0015
| U |  |  |  | 0.42 | 0.445 | -5.9524 | 0.002
Recommended values for international reference materials BE-N and BHVO-1 are from Govindaraju (1994, Geostandard Newsletter, Special Issue, v. 118, 158 p.). Abbreviations: Recomm. = recommended values; n.d. = not detected. Detection limits were estimated using 29 international reference standards and 3 internal standards run as unknowns.